Witness testimony for Element 4.1.2

“Measures and verifies optical appliances taking into account relevant standards where applicable”

Name of trainee: Date:

|  |
| --- |
| Initial box to indicate what has been completedFor each of 5 separate pairs of multifocal spectacles, I observed the trainee;* Measure and verify that lenses have been produced to the given prescription within B.S. tolerances.
* Verify that all aspects of the frame or mount have been correctly supplied
* Measure and verify that the lenses are corrected positioned within the spectacle frame/mount within B.S. tolerances.
* The trainee was able to answer my questions to demonstrate adequate understanding of the techniques and results obtained.
 |
| Additional guidance; Choice of instrumentation could include:* manual or semi-automated focimeter. (Fully automated e.g. Eye refract VX40 is not acceptable)

Accurate results to within:* ± 0.25DS/DC for dioptric measurements
* Axis appropriate to cylinder power
	+ ≤ 0.50DC ± 9°
	+ > 0.50DC ≤ 0.75DC ± 6 °
	+ > 0.75DC ≤ 1.50DC ± 4 °
	+ > 1.50DC ± 3 °
* Centres – 1mm tolerance.

Must demonstrate a knowledge of actual tolerances.BS EN ISO 21987:2017 |
| Brief summary of routine used (completed by trainee) |

Name of witness (block capitals) \_ Position of witness Witness GOC No.

Witness signature

Witness email and postal addresses